## Slide 1
ATLAS Forward Proton Time-of-Flight Detector LHC Run2 performance and experiences
Tomáš Sýkora
Institute of Particle and Nuclear Physics
Faculty of Mathematics and Physics
Charles University
on behalf of the ATLAS collaboration
Instrumentation for Colliding Beam Physics, Novosibirsk, 2020

## Slide 2: abstract
The performance of the ATLAS Forward Proton Time-of-Flight (ToF) Cherenkov detector is shown using the ATLAS collaboration data collected in the 2017 running period of the LHC Run2. The detailed analysis of the results, including detector efficiency and time resolution of the ToF is performed. The detector construction and its expected performance based on beam tests results are briefly summarized at the beginning of the first part of the talk. Operational experiences, problems caused due to attempts to operate Micro-Channel Plate PhotoMulTipliers (MCP-PMTs) in the vacuum and leading to the change of the detector construction, the request for R2D2 like construction of a new, future, MCP-PMT due to the observed gain drop for high rates and its non-recoverability are described. Also, other tested hardware changes are presented – the effect of the replacement of non-monolithic by monolithic crystal bars, the change of the MCP-PMT back-end electronics based on simulation to reduce signal cross talk resulting in further detector time resolution improvement. The second part of the talk is devoted to the achieved time resolutions, 20 ± 4 ps and 26 ± 5 ps, of two installed ToF detectors. Despite of their very low efficiencies (below 10%) in major parts of the analyzed data, this represents a superb time resolution for detectors operating a few millimeters from the LHC beams and making them, from the time resolution point of view, the best ToF detector among those operated by the different LHC experiments in the forward regions during the LHC Run2. At the end the possibility of reconstruction of z-coordinate of the ATLAS interaction region using the ToF detectors is shown.

## Slide 3: outline
AFP – physics and ToF subdetector
AFP ToF design in 2017
AFP ToF – expected performance
AFP ToF – operational experience in 2017
  MCP-PMT in vacuum
  crosstalk
  high rates
AFP ToF – new design
  Out-of-Vacuum solution
  LQBars
  trigger module (x)
  pulser (x)
AFP ToF – results
  timing
  z-vertex reconstruction
ATLAS-AFP Collaboration: Technical Design Report for the ATLAS Forward Proton Detector, CERN-LHCC-2015-009; ATLAS-TDR-024-2015
LANGE, J. et al., Beam tests of an integrated prototype of the ATLAS Forward Proton detector, JINST 11 (2016), no. 09, P09005
NOZKA, L. et al., Design of Cherenkov bars for the optical part of the time-of-flight detector in Geant4, Optics Express, 2014, vol 22(23), pp 28984-28996
NOZKA, L. et al. Construction of the optical part of a time-of-light detector prototype for the AFP detector, Optics Express, 2016, vol 24(24), pp 27951- 27960
CHYTKA, L. et al., Timing resolution studies ofthe optical part of the AFP Time-of-flight detector, Opt. Express. 2018, vol. 26, no. 7, pp. 8028–8039
CHYTKA, L. et al., Time resolution of the SiPM-NUV3S, NIM A, Volume 935, 11 August 2019
ČERNÝ, K., The ATLAS Forward Proton Time-of-Flight Detector System, International Workshop on Vertex Detectors 2019, Croatia
MELIKYAN, Y. at al., Load capacity and recovery behaviour of ALD-coated MCP-PMTs, NIM A, Volume 949, 1 January 2020, 162854
KOMÁREK, T. at al., Timing resolution and rate capability of Photonis miniPlanacon XPM85212/A1-S MCP-PMT, submitted to NIM A

## Slide 4: AFP – physics and ToF subdetector
AFP 217
AFP 205
AFP 205
AFP 217
beamline
beamline
ATLAS
p
p
a combination of AFP ToF information with e.g. HGTD…
ToF main goal – pileup suppression; possibility for a dedicated trigger

## Slide 5: AFP – location
ATLAS
Patch Panel 211 m
AFP stations
left-right (IP)
one side
XRP-NEAR
205 m
205m
AFP near
XRP-FAR 217 m
217m
AFP far

## Slide 6: AFP – design
PMT
CFD
HPTDC
FPGA ToF Trigger
2/4 bars/train
RF Switch
CTP
SiT Trig/IO HitOR-LTB
2/4 Planes
280 m
RCE-HSIO
320 m
Opto-Board
SiT Trig/IO HitOR-LTB
2/4 Planes
RG59 coax
AirCore coax
fiber
p
AFP station with passing proton p

## Slide 7: AFP ToF design in 2017
SiT
ToF
MCP-PMT
LQbars
SiT: 3D Pixels
50μm(x)×250μm(y)
σx≃7 μm/track
ToF: Cerenkov 4 Trains × 4 Bars
σt≃30 ps/Bar
resolution 10 ps or better for higher pileup 30 ps for 1-st phase
acceptance over full range of tracker
near 100% efficiency
segmentation for multi-proton timing
high rate capability, ~ 5 MHz/segment
L1 trigger capability
radiation hardness
MCP-PMT high rates capabilities
not possible with current design

## Slide 8: ToF performance as measured in beam tests
oscilloscope
measurements
the hunt for
final resolution after trigger subtraction – 20.6 ps
HPTDC included (RCE), trigger included
bars
PMT
PA-a + PA-b
PA-c
CFD
TDC
scope
π (SPS, 40/120 GeV)
RCE

## Slide 9: ToF components
contributions to time resolution
MCP-PMT < 10 ps (low rates, see below)
PAa (first layer of preamps)
PAb (second layer) 3-4 ps
CFD (constant fraction discriminator) 5 ps
HPTDC (high performance TDC) 15-17 (-25; can be corrected) ps
picoTDC < 6 ps
reference clock (synchronization) 6 ps (conservative)
time resolution limited by
bar design
MCP-PMT rate capability

## Slide 10: MCP-PMT – operability in vacuum
AFP ToF baseline
change of the detector design
repeated failures during HV tests in vacuum

## Slide 11: MCP-PMT – crosstalk ()
ceramic base
anode wire
anode pad
observed crosstalk, depending on the model (e.g. reduced anode gap): 20-40 %
motivated (also) by ALICE – change of back-end electronics design
modelling and simulations

## Slide 12: MCP-PMT – crosstalk modelling
signal and crosstalk – active 1 channel MCP-PMT
good agreement of simulation with measurement
signal and crosstalk – signal in 1-4 channels
new design, crosstalk: < 13 % for reduced anode gap

## Slide 13: MCP-PMT and high rates, Photonis, XPM85212

## Slide 14: MCP-PMT and high rates, Hamamatsu, R10754-07-M16

## Slide 15: MCP-PMT and high rates recovery
Gain variation with the increase of the average anode current of each device caused by a proportional increase in the illumination rate.
Device # | 1 | 2 | 3
MCP-PMT type | R10754-07-M16 | PMT253 | XP85002/FIT-Q
Manufacturer | Hamamatsu | Photek | Photonis
ALD-coating of the MCPs | yes | yes | no
Resistance of the MCP stack | 40 M? | 15.5 M | 14 M
Total bias voltage | 1600 V | 3000 V | 1285 V
Voltage across the MCP stack |  | 1704 V | 1018 V
Sensitive area of an individual readout channel | 5.3x5.3 mm2 | 6.6x6.6 mm2 | 26.5x26.5 mm2
Serial number | KT0714 | A1171005 | 9002140

## Slide 16: detector new design

## Slide 17: AFP insertions 2017
MD1-TS1 30.06-09.07
MD2 24-27.07
MD3-TS2 13-22.09
SpecRun-MD4 11-30.11
Xe run 12.10
ATLAS delivered/received ntegrated Luminosity
50
40
30
20
10
0
[fb–1]
Total Delivered to ATLAS: 50.4 fb–1
Total Recorded by ATLAS: 47.1 fb–1
Total AFP in Combined DAQ: 29.6 fb–1
Insertion: 12σ + 0.8 mm
Insertion: 12σ + 0.3 mm
Insertion: 11.5σ + 0.3 mm

## Slide 18: LQBars – glued vs monolithic solution
glued bars
– advantages
simpler and cheaper production – less groundwork (preparation phase, auxiliary elements), better machining (grinding, polishing)
excellent polishing quality up to 1 interference ring (peak-to-valey = 320 nm) on all surfaces
– disadvantages
glue – wavelength cut (235 nm) [?], additional tilt between arms, less radiation hard than used glass, fragility (small joint area, for bars with 2 mm radiator excessively small 2x5 mm2)
700 kGy ~ 100 fb-1 at 5 mm from beam
20 kGy ~ 100 fb-1 at 5 cm from beam

## Slide 19: LQBars – irradiated vs non irradiated, timing
bars irradiated by ~ 50.4 fb–1
bars measured during the AFP beam test in DESY 2019
no significant difference

## Slide 20: irradiated vs non-irradiated – efficiencies
non-irradiated bars
irradiated bars
1B
1C
1D
1B
1C
1D
no significant difference

## Slide 21: LQBars – glued vs monolithic - timing
monolithic bars give worse resolution!

## Slide 22: ToF in 2017 – timing
resolution of each bar within train and station, fully consistent with beam tests; no effect of synchronization clock

## Slide 23: ToF in 2017 – z-vertex reconstruction
zATLAS - zToF distribution: obtained by random mixing of times measured by the AFP To detector in either station and the zATLAS values which do not originate in the same collision event.
the expected resolution of the ToF detector, σToFexpected is obtained from the known single-channel resolutions convoluted with the actual channel-hit-patterns observed in the data in the 'no N_vtx cut' scenario
the shape of the distribution (narrow peak) provides a hint of presence of the (central diffraction) signal pp → pXp

## Slide 24: summary
AFP ToF detector was installed and took data in 2017
achieved timing performance was very good, despite of detector very low efficiency (which affected timing), in agreement with beam tests results
due to repeated failures of MCP-PMTs in vacuum, the detector was not installed in 2018
new design of detector ready
  MCP-PMT out of vacuum, tested in 2019
  new back-end electronics to suppress cross-talk, tested in 2020
  new sets of monolithic, glue-less, bars in production
  simplified signal and powering paths, will be tested in April 2020 beam test
  prototype of picoTDC tested in 2019, further tests planned in 2020
  new trigger module partially tested 2019; testing is ongoing, will be finalized during April’s beam test
detector will be installed for the LHC Run3 during summer 2020
open question – new type (R2D2) of MCP-PMTs [Arradiance + Photonis], answer within ~ 3 month from now
analysis of the 2017 data shown the good agreement with beam tests results and confirmed ToF detector potential to achieve resolution 15 ps in the LHC Run3 running

## Slide 25
backup slides