## Slide 1: The Upgrade of the ALICE TPC
Bernhard Ketzer
Universität Bonn
On behalf of the ALICE TPC Collaboration
26 February 2014
Instrumentation for Colliding Beam Physics 2014

## Slide 2: Time Projection Chamber
Combination of MWPC and Drift Chamber: 3-D tracking device
  long drift path (~m) in gas-filled volume  z coordinate
  MWPC + pads perpendicular to drift path  x,y coordinates
Time Projection Chamber [D.R. Nygren et al., Phys. Today 31, 46 (1978)]

## Slide 3: Advantages and Limitations of a TPC
An (almost) ideal tracking detector:
Large acceptance
Large active volume
Low material budget
3D hits  simple pattern recognition
Extremely high particle densities
 heavy ion experiments
Good momentum resolution
Particle identification
Examples:
STAR (420 cm × 400 cm)
ALICE (500 cm × 500 cm)
Limitations:
No redundancy
Calibration very demanding
Drift distortions due to ion backflow
Gating  low trigger rates

## Slide 4: Advantages and Limitations of a TPC
Examples:
STAR (420 cm × 400 cm)
ALICE (500 cm × 500 cm)
Limitations:
No redundancy
Calibration very demanding
Drift distortions due to ion backflow
Gating  low trigger rates

## Slide 5: The ALICE Detector
Volume: ~ 90m3
Gas: Ne/CO2(/N2) (90/10/5)
Drift voltage: 100kV, 94ms drift time
72 MWPCs with 557768 readout pads
Main tracking device: TPC
at 10 GeV
-0.9 < h < 0.9
(159 clusters)
Readout rate: 500 Hz (Pb-Pb)

## Slide 6: ALICE Upgrade
- Small diameter beam pipe

## Slide 7: Operation at High Rates
  Goal: Operate ALICE at high luminosity )  7 × more
  Record all minimum bias events
       50 kHz in Pb-Pb collisions, i.e. 100× higher than present
  ~ 5 overlapping events
  no gating and continuous readout: TPC as analog event pipeline
Event from
2013 p-Pb
Run

## Slide 8: Outline
Introduction
Ion Backflow in GEM Detectors
Prototype Detectors
  FOPI TPC
  ALICE Full Size Prototype
The TPC Upgrade
Expected Performance
Conclusions and Outlook

## Slide 9: Outline
Introduction
Ion Backflow in GEM Detectors
Prototype Detectors
  FOPI TPC
  ALICE Full Size Prototype
The TPC Upgrade
Expected Performance
Conclusions and Outlook

## Slide 10: Ion Backflow Suppression in GEM
Low ion density in drift region requires
  low primary ionization
  low gain
  low ion backflow

## Slide 11: The Space Charge Challenge
Ion blocking in GEMs not as efficient as with gating grid
Goal: IB = 1%, e = 20 at G = 2000
ions from 8000 events in drift volume!
Distortions up to 19 cm in and 7 cm in (near drift cathode for small radii)
Well below 10 cm for the largest part of drift volume
Similar to STAR TPC at RHIC
to be corrected to the level of intrinsic resolution!

## Slide 12: Ion Backflow – Early Measurements
Low ion backflow requires
    low drift field
    very asymmetric transfer fields
    highest amplification in last GEM
Ar/CO2 (70/30)
TUM
Edrift=200V/cm
UGEM1,2=310V
UGEM3=350V
Etrans1=6kV/cm
Etrans2=60V/cm
Eind=8kV/cm
Ar/CH4/CO2 (93/5/2)
IB ~ 1% easy to achieve!
Magnetic field helps:
  Electron extraction from holes improved

## Slide 13: Ion Backflow – Early Measurements
Low ion backflow requires
    low drift field
    very asymmetric transfer fields
    highest amplification in 3rd GEM
Ar/CO2 (70/30)
TUM
Edrift=200V/cm
UGEM1,2=310V
UGEM3=350V
Etrans1=6kV/cm
Etrans2=60V/cm
Eind=8kV/cm
Ar/CH4/CO2 (93/5/2)
Magnetic field helps:
  Electron extraction from holes improved
IB ~ 1% easy to achieve?

## Slide 14: Ion Backflow – Rate Dependence
Ar/CO2 (70/30)
Rate dependence of IB !
CERN

## Slide 15: Ion Backflow – Rate Dependence
Gain dependence of IB !
[A. Bondar et al., NIM A 496, 325 (2003)]

## Slide 16: Ion Backflow – Rate Dependence
Ar/CO2 (70/30)
Poisson equation:
For homogenous space charge density and parallel plate boundary conditions:
d
drift
z
Dependence of IB on drift gap!

## Slide 17: Ion Backflow – Rate Dependence
Ar/CO2 (70/30)
Poisson equation:
For homogenous space charge density and parallel plate boundary conditions:
d
drift
z
Expected after LS2: 5000 fC/cm2

## Slide 18: Ion Backflow – Triple GEM
Ne/CO2 (90/10)
Further issues to be considered:
Ne-based mixtures:
  Higher drift field
  a vs E steeper
 Add N2 to gas to avoid gap amplification
High electron collection efficiency (first GEM!)
Stability of operation
 Three GEMs not sufficient!
Add 4th GEM, try different aspect ratios:
standard (S) vs large-pitch (LP)
[F. Sauli et al., NIM A 560 (2006) 269]

## Slide 19: Ion Backflow – ALICE Solution
S-S-LP-S
S-LP-LP-S
Quadruple GEM stack, block ions by high-low combinations of fields
Use large-pitch GEMs to misalign holes
Optimize energy resolution at the same time
 IB = 0.6 – 0.8%, s = 11 – 12 % achieved

## Slide 20: Ion Backflow - Simulations
ANSYS + GARFIELD/MAGBOLTZ: Systematic scan of parameter space
    Gas composition
    Three / four GEMs, different geometries
    Tuning of simulations by adjusting hole alignment
3 GEM
4 GEM

## Slide 21: Outline
Introduction
Ion Backflow in GEM Detectors
Prototype Detectors
  FOPI TPC
  ALICE Full Size Prototype
The TPC Upgrade
Expected Performance
Conclusions and Outlook

## Slide 22: FOPI GEM-TPC
Light-weight field cage
Drift length 725 mm
Ø 105-300 mm
Triple GEM amplification
10254 ch., AFTER ASIC
Gas: Ar (Ne)/CO2 (90/10)
[L. Fabbietti et al., NIM A 628, 204 (2011),
M. Ball et al., arXiv1207.0013, 2012]

## Slide 23: dE/dx Resolution
Gain equalisation using 83mKr:
  Main peak 5.1%
PID using dE/dx:
  Resolution ~ 18%
  in agreement with expectation
  no density correction yet
First dE/dx measurement
with GEM-TPC!
[F.V. Böhmer et al., NIM A 737, 214 (2014)]

## Slide 24: ALICE Full Size Prototype
3 large-size GEM foils: single-mask
18 sectors (top side), ~ 100cm2 each
bias resistors 10 / 1MW
2mm frames glued on bottom side
spacer grid: 400mm thickness
additional frame for induction gap: 4mm

## Slide 25: dE/dx Resolution
Beam test at CERN PS: p, p±, e± beam, 1/2/3/6 GeV/c, LCTPC electronics
Gain equalization using tracks
No temperature/pressure correction yet
Truncated mean of cluster charge: 5 – 70% (only 46 pad rows considered)
for pions at 1 GeV
For comparison: IROC in ALICE TPC:
p=1 GeV/c
G=5000
[B. Ketzer et al., NIM A 732, 237 (1013)]

## Slide 26: Outline
Introduction
Ion Backflow in GEM Detectors
Prototype Detectors
  FOPI TPC
  ALICE Full Size Prototype
The TPC Upgrade
Expected Performance
Conclusions and Outlook

## Slide 27: Upgrade of the ALICE TPC
IROC
OROC
50cm
114cm
Single-mask technology
Quadruple GEM stack
Std + large-pitch foils
47cm
87cm

## Slide 28: Front-end Electronics
Requirements
Polarity opposite wrt MWPC
Continuous readout  simultaneous sampling and data transfer
Data throughput: (of FEE, not written to disk)
Triggered mode for calibration

## Slide 29: Outline
Introduction
Ion Backflow in GEM Detectors
Prototype Detectors
  FOPI TPC
  ALICE Full Size Prototype
The TPC Upgrade
Expected Performance
Conclusions and Outlook

## Slide 30: Expected Spatial Resolution
GEM pad response function smaller than MWPC  pad size adequate?
Resolution diffusion-dominated for , ~10% worse for GEMs
Cluster size ~20% smaller for GEMs  smaller occupancy
Increasing number of single-pad clusters for shorter drift length,
but partially outside barrel acceptance (|h|<0.9)

## Slide 31: Expected Momentum Resolution
Intrinsic performance:
GEMs slightly worse than MWPC due to space-point resolution
Recovered by vertex constraint + ITS
Performance with GEMs and pileup:
As good as present system with vertex constraint and ITS
Very little effect of pileup tracks (small PRF leads to less merged clusters)

## Slide 32: Reconstruction Strategy
Two-stage reconstruction scheme:
Cluster finding, cluster-to-track association: 1 TB/s  50 GB/s
  Scaled average space charge distortion map, , gain eq. 
Tracking, ITS-TRD track matching
  High-resolution space charge correction  in

## Slide 33: Reconstruction Strategy

## Slide 34: Conclusions
ALICE physics program for RUN3 requires an upgrade of the detector
TPC will be operated in continuous mode at 50 kHz interaction rate
GEM detectors instead of MWPCs to reduce the ion backflow
  4 GEMs with different aspect ratios needed to meet requirements
  IB < 1%, energy resolution s < 12% achieved
  Room to optimize operational stability
Distortion corrections applied in two steps
  online using average map  resolution
  offline using high-resolution space charge map (~ ms) 
Performance of TPC can be fully maintained at 50 kHz
TDR ready for submission

## Slide 35
Thank you for your attention!