## Slide 1: An APD sensor with extended UV response for readout of BaF2 scintillating crystals David Hitlin INSTR2014 February 28, 2014

## Slide 2: Background and motivation
When the cost of LYSO reached unaffordable levels, Mu2e needed a fast, radiation hard crystal with reasonable light output
Barium fluoride is a potentially interesting candidate, but it presents unique problems
  It has a very fast scintillation component, which is attractive, but it is accompanied by a much larger slow component
  Both components are in the UV
Making use of BaF2 in a practical experiment requires a photosensor that
      discriminates between fast and slow scintillation components
      has a fast time response
      works in a magnetic field
      is stable over time
      is radiation hard
The development of such a sensor is the subject of this talk

## Slide 3
Fast scintillating crystals
Ren-yuan Zhu

## Slide 4: Scintillation pulse shapes
BaF2
Ren-yuan Zhu
LaBr3
LSO
LYSO
LaCl3

## Slide 5
Ren-yuan Zhu
A fast crystal “figure of merit”
Motivates R&D on fast crystals and appropriate solid state readout

## Slide 6: BaF2 is a potentially attractive high rate crystal
BaF2 is among the fastest scintillating crystals (0.9ns), but there is a much larger, slower, component (650ns)
In order to take full advantage of the fast component, it is necessary to suppress the slow component
85%
650 ns
15%
900 ps
Total light output 1.2 x 104 photons/MeV

## Slide 7: BaF2 is a potentially attractive high rate crystal
BaF2 is among the fastest scintillating crystals (0.9ns), but there is a much larger, slower, component (650ns)
In order to take full advantage of the fast component, it is necessary to suppress the slow component:
  Need a “solar-blind” photosensor

## Slide 8: BaF2 is a potentially attractive high rate crystal
BaF2 is among the fastest scintillating crystals (0.9ns), but there is a much larger, slower, component (650ns)
In order to take full advantage of the fast component, it is necessary to suppress the slow component
  La doping of pure BaF2 suppresses the slow component by ~4
  Other dopings can be explored

## Slide 9: Solar-blind photosensors
Solar-blind PMTs exist
  Large area, fast, but expensive, and do not work in a magnetic field
Solar-blind solid state devices also exist
  SiC APDs (100mm diameter)
  AlGaN APDs (< 1mm diameter)
There are several potential approaches to fast, large area, solar-blind, magnetic field insensitive photosensors
  A variant of the LAPPD channel plate under development by U Chicago/Argonne
  SiPMs with or without antireflection coatings (e.g., Hamamatsu)
  Large area delta-doped APDs with AR ALD (Caltech/JPL/RMD)

## Slide 10: How to achieve best possible QE in the 200-300 nm regime?
Absorption length at 220 nm in silicon is less than 10 nm
  Protective epoxy coating on an APD or SiPM has a strong effect on QE
  Sensitive region of device must be very close to the surface

## Slide 11: UV sensitive MPPC
Requirements
  Sensitivity to liquid xenon scintillation (λ = 175 nm)
  Large active area (12×12 mm2)
  Single photon counting capability
  Moderate trailing time constant (τ ＜ 50 ns)
In order to improve sensitivity
  Remove protection layer
  Match refractive index
  to liquid xenon
  Apply anti-reflection coating
Hamamatsu Photonics/MEG
Cross-sectional image of MPPC

## Slide 12: IEEE/NSS Seoul
Large area MPPC sensitive to the liquid xenon scintillation light ( λ = 175 nm) 12×12 mm active area Detection efficiency (PDE) of 17% Pixel gain around 106 Version with improved pixel structure is under development
Performance of UV-Sensitive MPPC for Liquid Xenon Detector in MEG Experiment
D. Kaneko
ICEPP, The University of Tokyo, Tokyo, Japan
On behalf of the MEG Collaboration
Hamamatsu UV sensitive MPPC
P. Murat

## Slide 13
Far ultraviolet, visible-blind coatings
Silicon
HfO2
Al2O3
Native SiO2
World record deep ultraviolet quantum efficiency
  QE > 5x Galex
  Stable
  Visible blind coatings are possible
“Ultraviolet antireflection coatings for use in silicon detector design,”
Erika T. Hamden, Frank Greer, Michael E. Hoenk, Jordana Blacksberg, Matthew R. Dickie, Shouleh Nikzad, D. Christopher Martin, and David Schiminovich
Applied Optics, Vol. 50, Issue 21, pp. 4180-4188 (2011)
“Delta-doped electron-multiplied CCD with absolute quantum efficiency over 50% in the near to far ultraviolet range for single photon counting applications”
Shouleh Nikzad, Michael E. Hoenk, Frank Greer, Blake Jacquot, Steve Monacos, Todd J. Jones, Jordana Blacksberg, Erika Hamden, David Schiminovich, Chris Martin, and Patrick Morrissey
Applied Optics, Vol. 51, Issue 3, pp. 365-369 (2012)
“Atomically precise surface engineering of silicon CCDs for
enhanced UV quantum efficiency,”
Frank Greer, Erika Hamden, Blake C. Jacquot, Michael E. Hoenk, Todd J. Jones, Matthew R. Dickie, Steve P. Monacos, Shouleh Nikzad
J. Vac. Sci. Technol., A 31, 01A103 (2013)

## Slide 14: APD structure
Deep-diffused architecture (RMD)
Photoelectrons created by UV photons at the sensing surface must survive trapping at the SiO2/Si passivation surface interface and recombination in the undepleted p-side neutral drift region (tens of mm) to reach the depletion region where the avalanche takes place
  The thickness of the undepleted p-side region is engineered for efficient conversion of visible photons
    This is undesirable for high UV QE
The result is that UV QE is reduced and the device speed is determined by the ~10 ns drift time as well as by capacitance
This structure can be modified, using proven techniques, to improve both UV quantum efficiency and device timing characteristics

## Slide 15: Visible-blind, superlattice-doped silicon detectors
  Eliminating the potential well at the surface and reducing (or removing) the drift region improves the detector UV QE and speed
  The process is executed on modified RMD APD wafers, using an extrapolation of existing technology at Caltech’s Jet Propulsion Laboratory (JPL)
    Molecular Beam Epitaxy
      > 50% QE at 220 nm
      Low dark current
      High conductivity
      Radiation tolerance
    Atomic Layer Deposition
      Chemically passivated interface
      Visible-blind antireflection coatings: < 1% QE at 300 nm
RMD 9x9mm

## Slide 16: The process
  Wafers are thinned on the p-side (sensing side) to remove the undepleted region
  Drift region is replaced by 3mm vapor phase epitaxial layer
  Monolayer of boron atoms is deposited ~2.5nm below the surface, eliminating the trapping well
MBE growth
3 nm
Delta-doped layer
(dopant in single atomic layer)
Native oxide
1 nm
original Si
circuitry

## Slide 17: Delta-doping has been applied to CCD detectors
S. Nikzad, “Ultrastable and uniform EUV and UV detectors,” SPIE Proc., Vol. 4139, pp. 250-258 (2000).
Thinned CMOS Array
Back surface
front surface
J. Trauger (PI WF/PC2) – No measurable hysteresis in delta-doped CCDs

## Slide 18
Bulk
Surface
Positively charged surface @ 1013 cm-2
Delta doping and quantum exclusion
  Delta-doping creates a “quantum well” in the silicon
  Majority carriers are confined to quantized subbands
  Quantum exclusion eliminates trapping Peak electric field is ~107 V/cm

## Slide 19
Superlattice doped surface with delta doping
5 nm, 8x1014 cm-2
2.5 nm, 2x1014 cm-2

## Slide 20: Thinning improves timing performance
3 mm epitaxial layer deposited on thinned RMD APD (3x3 mm2)
Illumination with 405 nm laser pulse

## Slide 21: Experimental tests for Al2O3/Al Blocking Filters (220 nm)
Design Approach:
  Develop optical model for stand-alone Al2O3 layers (these were deposited at 200°C, Oxford ALD w/ O2 plasma)
  Using this model and Al reference model (Palik data) calculate for target layer thickness
  Designs roughly based on original rejection targets for three or five layer metal/dielectric stacks (target peak @220 nm)
  Feedback from ellipsometry data to modify thickness targets (i.e. due to Al layer oxidation duration processing, etc.)
  Will eventually develop a refined optical model to be more predictive about the expected transmission and reflection, simple discrete layers appear insufficient especially for the five layer tests
J. Hennessy

## Slide 22: Five Layer Test (Si-DMDMD)
Five layer w/ new Al parameters
Best fit χ2 = 9.61 (L1 = 418Å, L2=217Å, L3=392Å, L4=158Å, L5=828 Å)
Model T @220nm = 38% (45°) , 53% (normal)
Model T @310nm = 0.16% (45°) , 0.20% (normal)
Predicted Tpeak (normal) = 55% @ 218 nm
Five layer w/ ideal Al parameters
Best fit χ2= 25.6 (L1 = 426Å, L2=204Å, L3=400Å, L4=161Å, L5=828Å)
Model T @220nm = 53% (45°) , 68% (normal)
Model T @310nm = 0.13% (45°) , 0.16% (normal)
Predicted Tpeak (normal) = 69% @ 218 nm
Solid Red = model transmission
Solid Blue = model reflection
Models based on best-fit to ellipsometric data
Symbols = measured reflection
*All 45° incidence
J. Hennessy

## Slide 23: Angular dependence of integrated interference filter
J. Hennessy

## Slide 24
  RMD has furnished six thinned APD wafers to JPL for processing
      Two stage development plan
        I: Extended UV response
        II: Integrate AR interference filter
  Processed wafers have been probed and diced at RMD
    Devices work as APDs, with somewhat increased noise but fast pulse response
      Noise has been traced to a contaminated epitaxial vapor deposition source at JPL, which is being disassembled and cleaned
    JPL will process and AR coat several unpackaged chips and return both coated and uncoated chips to RMD for packaging
    Chips will undergo full characterization and then tests with BaF2
  A second batch of wafers is being readied
    Will explore different depths and density of the superlattice MBE and delta doping
  Goal is to have full size prototype devices for a beam test at MAMI in the fall
Current status

## Slide 25: Conclusions
Full utilization on the fast component of BaF2 requires development of an appropriate sensor
Existing large area UV sensitive APDs or SiPMs have ~20% quantum efficiency, slow time response and offer no discrimination between the BaF2 fast and slow scintillation components
The Caltech/JPL/RMD collaboration is developing a 9x9mm superlattice thinned, delta-doped APD with integrated AR coating and filter
  This device promises to have >50% quantum efficiency at 220 nm, strong discrimination of the 300 nm slow component, and excellent timing characteristics

## Slide 26: The proverbial bottom line

## Slide 28: Scintillating crystal calorimeter for mu2e
Two disk geometry
Hexagonal BaF2 crystals; APD or SiPM readout
Provides precise timing, PID, background rejection, alternate track seed and possible calibration trigger.

## Slide 29: Calorimeter crystal history
Initial choice PbWO4: small X0, low light yield, low temperature operation, temperature and rate dependence of light output
CDR choice LYSO: small X0, high light yield, expensive (→very expensive)
TDR choice: BaF2: larger X0, lower light yield (in the UV), very fast component at 220 nm, readout R&D required, cheaper,

## Slide 30: Reach-Through Avalanche Photodiode (RTAPD)
N+
contact
p+ implant
pn+ junction
Reverse biased photodiode with p+πpn+ structure
Absorption
Multiplication
e-
Recombination
silicon
Drift: velocity 10 ps/μm
e-
e-
e-
e-
P+
P+
Ultraviolet light is absorbed near the surface
Electric field drives electrons toward junction
Avalanche gain in pn+ junction
Dead layer formed by p+ implant
e-
Signal
220 nm
Deep implant UV QE Low, unstable
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
Background
330 nm
No background rejection

## Slide 31: Atomic Layer Deposition (ALD)
Achieves layer-by-layer growth of films with Angstrom-level control over arbitrarily large surface areas
Wide suite of materials metals, oxides, and nitrides with excellent film properties
Can be directly integrated into existing detectors/instruments to vastly improve performance
TEM images of ultra-thin (3.5nm), conformal ALD film
Key advantages of ALD
Fully complements the surface engineering capabilities of Si MBE with atomic control of ALD
Thickness can be specified with Angstrom resolution
  Enables precise, repeatable targeting of bands e.g. 16.5 nm vs. 23 nm of Al2O3
Process is completely independent of device size
3.5 nm
ALD
Silicon Substrate
Silicon Substrate
3.5 nm
ALD
Glue for sample mount